## Slide 1: First results of the Channeling Exercise Status June 2018
Presenter: Thomas Blumenstock
with contributions by many of you!
Notes: Alle Beteiligten aufführen!

## Slide 2
AeMet Izana | O. Garcia | 
Bremen U. | M. Palm | J. Notholt
Gothenborg U. | G. Persson | J. Mellqvist
FMI, Sodankylä | P. Heikkinen | R. Kivi
KIT Garmisch | M. Rettinger | R. Sussmann
KIT Karlsruhe | F. Hase | T. Blumenstock
Liege U. | C. Servais | M. Mahieu
Nagoya U. | T. Naghama | 
NCAR Boulder | I. Ortega | J. Hannigan
NIES, Tsukuba | I. Morino | 
NIWA, NZ | J. Robinson | D. Smale
St. Petersburg | M. Makarova | 
Toronto U. | E. Lutsch, O. Colebatch | K. Strong
UNAM Mexico | W. Stremme | M. Grutter
Wollongong U. | N. Jones | D. Griffith
Contributions from
Thank you for
your effort!

## Slide 3
Motivation
Experiment
Results
Discussion
Outline
Taken from wikipedia.org

## Slide 4
Detector non-linearity (Abrams et al., 1994)
ILS (Instrumental Line Shape) (Hase et al., 1999)
Intensity fluctuations, DC correction (Keppel-Aleks, 2007)
Precise solar tracking (Gisi et al., 2011)
Sampling error (Messerschmidt et al., 2010; Dohe et al., 2013)
=> Channeling
Motivation
Previous work on error sources

## Slide 5
Talk at IRWG 2004 & 2008
Does not affect total columns of strong absorbers,
but leading error for weak signatures.
NDACC standard species only ClONO2, HNO3
Might affect profile retrieval of strong absorbers.
Recently work on species with weak signatures:
HCHO, SF6?, ....
Triggered by work on HCHO by Corinne & discussion
on error calculation by Frank
Vigouroux, C. et al.: NDACC harmonized formaldehyde time-series from 21 FTIR stations covering a wide range of column abundances, Atmos. Meas. Tech. Discuss., https://www.atmos-meas-tech-discuss.net/amt-2018-22/, in review, 2018.
Motivation
Why not addressed earlier?

## Slide 6
Vigouroux et al., AMTD, 2018
NDACC harmonized formaldehyde time-series
Retrieval scheme harmonised,
error calculation not yet:
PROFFIT includes channeling error, SFIT-4 not (yet).

## Slide 7
Lab source without glass body if available:
Globar or black body
NDACC filters #3 and #6 (or #7/8)
Spectral resolution of 0.05 cm-1
Co-add 1000 scans for a good signal to noise ratio
Outline of test
BB source
Kiruna setup

## Slide 8
Btw, transmission of filter #3

## Slide 9
Cut spectrum to approx. 50 cm-1
Normalize to 1
FFT
Analysis of test spectra
=> shows two dominant frequencies,
however absolute numbers needs correction!
=> Analysed manually!

## Slide 10
 | F 1 [cm-1] | A 1 [‰] | F 2 [cm-1] | A 2 [‰] | Comment
Eureka | 0.93 | 0.3 | 0.1 | 0.1 | 
Ny-Alesund | 0.91 | 5 |  |  | 
Kiruna | 0.9 | <0.3 |  |  | Lab air
Sodankylä | 0.93 | 1.0 | 0.1 | weak | F2 ~ noise
St.Petersbg. | 0.90 | 1.0 | 0.1 | 0.6 | 
Bremen | 0.93 | 0.8 |  | < 0.2 | F2 ~ noise
Garmisch | 0.91 | 1.2 | 0.1 | < 0.1 | F2 ~ noise
Zugspitze | 0.93 | 0.5 | 0.1 | 0.2 | 
JJoch | 0.92 | 2.8 |  |  | 
Toronto | 3.0 | 2.0 | 0.075 | ~ noise | BOMEM
Rikubetsu | 0.91 | 2.8 | 0.11 | 0.6 | 
Tsukuba | 0.92 | 2.8 | 0.1 | 1.0 | 
Izana | 0.9 | 0.3 | 0.1 | 0.4 | Lab air
M. Loa | 0.93 | 2.8 |  |  | 
Altzomoni | 0.9 | 0.3 | 0.1 | 0.2 | noise?
Results

## Slide 11
 | F 1 [cm-1] | A 1 [‰] | F 2 [cm-1] | A 2 [‰] | Comment
Wollongong | 0.93 | 1.0 | ~0.1 | 0.2 | 
Laud HR120 | 0.90 | 1.0 |  |  | 
Laud HR125 | 0.91 | 3.3 | 0.11 | 1.3 | 
A. Heights | 0.91 | 2.0 | 0.1 | 0.4 | 
Most spectrometers show 2 dominant channeling frequencies: 0.9 & 0.1 cm-1!
Frequencies fit to beamsplitter substrate and gap.
Amplitude (peak to peak) is up to 5 ‰, typically it is 1-3 ‰!
Amplitude is ~0.5 ‰ or lower for some sites:
Eureka, Kiruna, Zugspitze, Izana & Altzomoni.
Originally PROFFIT error calculation assumed 4 ‰ (pp) (FH)
For the paper channeling ampl. was set to 1 ‰ (pp) in total: Freq.: (0.005), 0.2, 1.0, 3.0 cm-1; 50% random/ 50% systematic
Results, ctd.

## Slide 12
KBr – BS wedges channeling frequencies:
Original: Sub.: 1‘, Gap ?, small 0.88 & 0.23 cm-1
New: Sub.: 8‘, Gap ? 0.88 cm-1
New 2: Sub.: 8‘, Gap ~1° __
      Dk = 1/(2ndcosQ):
      with n=1, d=5.7 mm: Dk = 0.88 cm-1
      with n=1.5, d=14 mm: Dk = 0.23 cm-1
Signatures of lab air in BB spectra: CO2, H2O
Much of ‘noise’ in atmospheric MCT spectra is due to channeling!
Larger wedge important for NDSC-IRWG group
No compatibility with FIR BS is the price to pay
CaF2 with 10’ is compatible
Alignment & ILS is still fine
Comparison of different wedges of KBr beam splitters
BUG 2004
1(2*1.5*3cm) = 0.11 cm-1

## Slide 13
Comparison of BB spectra with different KBr beam splitters
KBr – BS:
Sub. 8‘,
Gap ~1°
Substrate: 8‘
Original:
Sub. 1‘
BUG 2004

## Slide 14
Presentation by Arno Simon
Experience with upgrades to 125 HR/M
Channeling / beam splitter with bigger wedge
    Channeling frequencies: 0.88 & 0.23 cm-1 (see BUG 2004)
    KBr with 8’ / CaF2 with 10’
    New alignment of spectrometer is needed!
    No compatibility with FIR BS!
    Joint order!?
Any other issues?
Please feel free to use: ndacc-bug@acd.ucar.edu
Bruker User Group (BUG) Meeting 2008
BUG 2008

## Slide 15
Vigouroux et al., AMTD, 2018
TR<D30
TR<D30
TR<D30
TR<D30
*** TR>D30

## Slide 16
Conclusions
Since wedge of beamsplitter and its gap is too small most spectrometers show two dominant channeling frequencies: 0.9 & 0.1 cm-1!
Amplitude (peak to peak) is up to 5‰, typically it is 1-3 ‰!
Amplitude is 0.5 ‰ or lower for some sites:
Eureka, Kiruna, Zugspitze, Izana & Altzomoni.
Uniform channeling amplitude for error est. doesn’t fit.
Random error > variability for the sites with low channeling.
For some sites random error < observed variability:
Ny-Alesund, Toronto, Paramaribo, Wollongong & Lauder
Channeling error needs to be included in error estimate of HCHO column amounts!
Probably also for other weak absorbers as ClONO2, SF6 ...!
For some sites random error ~ variability: random smoothing
error (too large) compensates for missing channeling error?

## Slide 17
Continuing this exercise? E.g. on MCT spectra?
  Missing sites?
Beam splitters with bigger wedge in the future!?
To be discussed with Axel Keens, Bruker
Channeling error in error calculation is needed for HCHO!
Probably also for other weak absorbers as ClONO2, SF6 ...!
Will channeling be included in future SFIT-4 err. estimate?
Documentation (when finished)?
Topics to be discussed

## Slide 18
and participation!