## Slide 1
NIWA MIR-FTS program: Lauder & Arrival Heights site reports
Arrival Heights: 78S, 167E, 220m
Lauder: 45S, 169E, 370m
Dan Smale, John Robinson, Dave Pollard, Alex Geddes
Sanil Lad (AntNZ), Anitra Fraser (AntNZ)

## Slide 2
Summary
There is long-term, stable funding of both Lauder and Arrival Heights (AHTS) MIR-FTS measurements (~1.2FTE) with continuing logistical support from Antarctica New Zealand.
NDACC archiving of 10 species current to the end of 2017.
Nine out of 10 retrievals using SFIT4 and NDACC compliant retrieval strategies.
Profiles submitted, but total column only for pre-2001 Lauder (Bruker 120M) datasets.
Both the Lauder 120HR and AHTS 125HR have been operational over the past year with no prolonged downtime. Both Brukers exhibit stable ILS.
In May 2018, the main Lauder MIR-FTS instrument (NIWA001) was switched to a 125HR (NIWA006). A 6 month intercomparison (Nov 17-May18) between the 120HR and 125HR was completed.
We are continuing to replace/upgrade obsolete hardware and software to avoid any large gaps in the spectra time series: New AHTS tracker, new scheduling s/w and a new Lauder 125HR.
Continue to participate in IRWG network wide activities and aiming to supply fully-compliant data files to the NDACC archive in the coming year.

## Slide 3
NDACC submission
SFIT4 & NDACC-compliant retrievals. HDF format.
(IRWG Uniform Retrieval Parameter Summary, Hannigan, 2012, updated 2014 + species dependent strategy updates from IRWG members).
Only ClONO2 using SFIT2 (total column only, no error analysis), rest SFIT4 (9 species)
Lauder submissions, all 10 species:
Bruker 120M : 1990 – Oct 2001: Total column only.
Bruker 120HR: Oct 2001 – Dec 2016: Profiles
AHTS submissions, all 10 species:
Bomem DA2 (1992-1995): HNO3 and HCl (total column).
Bruker 120M (1996- Feb 2016): profile, (except O3, total column at the moment.)
Bruker 125HR (Dec 2014 – Dec 2016): profile. (except O3, total column at the moment.)
Error analysis:
Full timeseries: HF, HCl, HCN, N2O
Partial timeseries: CO, O3, CH4 (2016 onwards)
To do: C2H6, ClONO2, HNO3
Other species available on request: H2CO (Lauder only), NO2 & OCS
CAMS27: monthly submissions of CO, CH4 and O3 from both sites.

## Slide 4
Instrument performance over the past year: Lauder 120HR ILS HBr cell
Lauder120HR: ‘Noisy’ InSb Detector (high RMS, zero level erratic) in Sept 2017. A pump down fixed it.

## Slide 5
Instrument performance over the past year: AHTS 125HR ILS N2O cell
AHTS120HR: working fine, no issues.

## Slide 6
Observation days:

## Slide 7
A lesson learnt…..:
Caution should be taken when cleaning 120/5HR vacuum seals with Isopropyl alcohol (2-propanol: C3H8O). Ventilation is imperative.
Isopropyl alcohol absorption seen in N2O cell spectra:
deweighting regions for each MW
$
1 (2183.1,2183.4),0.001
0
0
C3H8O NIST reference spectra
res: 0.125cm-1
(scaled for clarity)
AHTS 125HR background spectra 0.0035cm-1. 50 scans, ~1000hpa
Prior to clean
After clean with C3H8O

## Slide 8
Hardware: A Bruker crèche at Lauder.
Previously two Brukers at Lauder:
MIR 120HR (NIWA001) - NDACC
NIR 125HR (NIWA006) – TCCON
A new 125HR arrived 20th May 2018 (NIWA008) to replace the 120HR.
NIWA001, NIWA006 and NIWA008 are all MIR/NIR capable.
We have the infrastructure (trackers + real estate) for 3 solar observing Brukers.
 | Optics building site#1 NIWA001, then 008 | Optics building site#2 NIWA006 | IDL building Site #3
Current set up | 120HR MIR (NDACC) Occasional NIR obs | 125HR NIR (TCCON) | AHTS 120M
Nov 2017 – May 2018 | 120HR MIR | 125HR NIR MIR comparison days | AHTS 120M*
New 125HR arrives | New 125HR NIR | 125HR MIR/NIR NIR comparison days | 120HR MIR
Final configuration | New 125HR NIR (TCCON) Occasional MIR obs | 125HR MIR (NDACC) Occasional NIR obs | 120HR Role?
This method allows intercomparison of both MIR and NIR measurements
on both instruments.
Overall a change in the NDACC & TCCON primary instruments.
(NIR and MIR inter comparisons required)
*Surplus to requirement at Lauder: 2 x 120M’s . These have been gifted to BIRA
NIWA001 and NIWA006 MIR intercomparison measurements.
From Nov 17 to May 2018, ~6 months (half the seasonal cycle)
6 NDACC filters, ~9 days per filter,
~100 (125HR) & 60 (120HR) spectra per filter
Analysis pending

## Slide 9
Technical transfer:
A Lauder first author paper:
Python-based dynamic scheduling assistant for atmospheric measurements by Bruker instruments using OPUS. Alexander Geddes, John Robinson, and Dan Smale Appl. Opt. 57(4), 689-691 (2018). (a technical note)
Lead by IRWG collegaues:
Olsen et al, AMT, 2017 : CH4 TANSO-FTS
Dammers et al, AMT, 2017: NH3 IASI
Ozone:
Steinbrecht et al, ACP, 2017 :An update on ozone profile trends for the period 2000 to 2016
note: Lauder the only site with lidar, microwave, umkehr and FTIR gb measurements.
Gaudel A, et al, Elem Sci Anth, 2018, Tropospheric Ozone Assessment Report: Present-day distribution and trends of tropospheric ozone relevant to climate and global atmospheric chemistry model evaluation
Fig 1, Steinbrecht, ACP, 2017

## Slide 10
Opportunistic measurement of Erebus Volcanic plume composition from solar occultation spectrometry:
Twice a year around noon (~4th Apr, ~5th Sept, SZA ~85), the sun skims the top of Erebus, Ross Island Antarctica.
Measurements thru the plume are possible from Arrival Heights if conditions hold. (only 4 days from 20+ years of obs)
Large HCl and HF enhancements seen. Plume SO2, SiF4, H2O?, CO2 minimal.
Dan, Jim H & Nicholas J working on a paper.
HCl 5/4/18
HF 5/4/18

## Slide 11
Thanks:
And lastly a big thank you to the following people who have provided invaluable assistance to us at Lauder:
Nicholas J, Jim H, Frank, Bavo, Mathias P, Corinne
Roger Lin (NDACC archiving)
Antarctica New Zealand
Gregor Surawicz (Bruker)
ATom overpass: 4th May 2018: a brilliant day
Photos: Dave Pollard, Dave Fedors