## Slide 1: Solar FTS measurements with a fibre optic coupled suntracker and EM27 FTS Side by side comparisons with TCCON, EM27SUN, IRcube and Aircore
David Griffith, Nicholas Jones, Nicholas Deutscher, Clare Murphy
University of Wollongong, Australia
Dave Pollard
NIWA, New Zealand
David Noone
University of Auckland, New Zealand

## Slide 2: Introduction & motivation
The EM27SUN requires weather protection. This requires either
  Manual daily set up, take indoors at night, OR
  Automated weather enclosure
A weatherproof tracker outdoors with fibre optic coupling to FTS located indoors largely avoids this
  And is nicer to the FTS, kept indoors . Better longevity
  Tested in FRM4GHG and before with Bruker IRcube and Ekotracker
Here we describe an EM27 coupled to a solar telescope on a commercial suntracker with 20 m fibre optic coupling
  Hardware setup
  Performance tests
  Side by side comparisons with TCCON, EM27 SUN, IRcube and Aircore

## Slide 3: History: IRcube and FO-Ekotracker FRM4GHG, Sodankyla, 2017, 2018

## Slide 4: Solar fibre-optic coupling to portable low-resolution FTSs EM27 and IRcube David Griffith, Nicholas Jones, Nicholas Deutscher University of Wollongong contact: griffith@uow.edu.au
Summary and Introduction
We present an alternative solar tracker setup to the conventional alt-azimuth 2-mirror tracker such as the Bruker / EM27SUN camtracker
  The sun is imaged into an optical fibre (FO) by a simple glass lens telescope mounted on a commercially-available Aeronet-style tracker
  The FO brings the solar radiation into the lab
  The divergent FO exit beam is collimated into the FTS for spectrum measurement
  EM27 spectra using the telescope-FO and Camtracker are compared
This setup allows the FTS to be operated indoors, avoiding harsh environments
We have successfully used the telescope-FO tracker system for several years in the FRM4GHG project with a Bruker IRcube FTS
Here we combine the telescope-FO input with an EM27 FTS and compare to EM27SUN with camtracker
Next steps …
Optimise fibre optic for robustness and reproducibility
Check for fibre spectral effects
  spectral structures, fringing
  Dependence on fibre bending and movement
  Not apparent so far
Routine data collection and analysis (GGG, PROFFAST)
Comparisons with co-located TCCON
  New HR125 at UoW
Comparisons with co-located COCCON
On the roof
Telescope  50mm x FL 35/40 cm
Fibre optics to EM27 or IRcube
Ekotracker (Japan)
Sodankyla
UoW
Conclusions
Spectra via the telescope-FO and the Camtracker are so far very comparable
Encourages further investigation to assess accuracy and variability
Details
Ekotracker: autonomous, active tracking <0.1°, self-parks at night. Robust, reliable, weatherproof
Telescope: BK7 glass lens,  50mm x FL 35 or 40 cm
Glass fibre optics 500 or 800 m, 20 m long, divergent exit beam ~ F/3
IRcube 0.5 cm1 resolution, single sided igm
EM27 0.5 cm-1 resolution, double sided igm
In the lab
IRcube at Sodankyla
FO input replaces camtracker
Camtracker view
Detector view
Comparative spectra
___ EM27 - Camtracker
___ EM27 - Fibre optic
___ EM27 - Camtracker
___ EM27 - Fibre optic
Fibre from roof
Collimating lens
EM27 at UoW
Pros
Tracker is autonomous, robust, reliable and weatherproof
FTS is housed indoors in protected environment
  Weatherproofing not required
Continuous 24/7 measurements
Comparable signal levels and SNR
Cons
Possible fibre spectral structures and variability
  Continuum shape, fringing
Fibre/glass cutoff below 4500 cm-1
  CO, N2O bands not accessible
Acknowledgements
We thank ESA and the FRM4GHG project team for their support for this work.
TCCON 2022

## Slide 5: EM27 with Ekotracker and fibre optic feed 1. Schematic
BK7 glass lens
Telescope: FL 350/400 mm, diam 50 mm
Fibre optic to lab
To EM27
Fibre optic from roof
25 mm lens
Solar measurement
Lamp measurement
To EM27
25 mm lens
Tracker

## Slide 6: EM27 with Ekotracker and fibre optic feed 2. On the roof – the tracker
Ekotracker, Japan
  Nominal tracking accuracy <0.01 (~1/50 solar diam)
  GPS + active tracking
  Fully autonomous
    Auto start/stop at night
  Serial command interface
  <$20,000
  Very robust and reliable
  No weather protection required
Fibre optic
  Quartz, multimode
  diameter 0.5 mm
  length 20 m
  Eg Thorlabs,
Fibre optics to lab EM27 and IRcube
350/400 mm FL telescopes
Ekotracker
Active tracking sensor

## Slide 7: EM27 with Ekotracker and fibre optic feed 3. In the lab – the FTS
Fibre optic from roof
Folding mirror
Collimating lens
Focussing lens
for lamp
Collecting mirror from lamp for ILS
Bolts to camtracker mounts

## Slide 8: EM27 with Ekotracker and fibre optic feed 4. Input alignment
Detector view
Camtracker view
FO end is imaged onto
0.6 mm detector aperture
The camtracker camera is useful to check input alignment
(but it is not actively tracked here)

## Slide 9: EM27-FO-Ekotracker Assessment
Spectra compared
Comparison to co-located TCCON (Wollongong)
  All EM27 data processed with GGG
  ADCF and AICF factors from Nasrin Pak et al. applied
Tracker alignment
  Solar-terrestrial line shift (O2_SG)
Spectral residuals
  Evidence for fringing or other FO artefacts?
Sensitivity to input alignment
  Fibre rotation, focus, intensity?
ILS

## Slide 10: Comparison spectra camtracker vs fibre optic
___ ___ EM27 – camtracker
___ ___ EM27 – Fibre optic
Fibre and lens cutoff

## Slide 11: Comparison spectra O2 band
___ EM27 – camtracker
___ EM27 – Fibre optic

## Slide 12: EM27-FO vs TCCON GGG analysis Xluft, XCO2, XCH4 => AICF
Xluft
XCO2
XCH4

## Slide 13: Check ADCF TCCON & EM27-FO vs SZA
Timing error?
Tracker error?

## Slide 14: Tracker pointing: O2_7885_sg Solar-terrestrial line shifts, same tracker, different telescopes
EM27-FO
IR cube

## Slide 15: Spectral residuals Is there evidence for fibre-induced fringing effects?

## Slide 16: Spectral residuals Is there evidence for fringing/interference effects from the FO?
Residual
FFT of Residual
Max OPD

## Slide 17: Side by side measurements Sept 2022 – May 2023
Instrument | Broken Hill / Aircore Sept - Oct 2022 |  | Wollongong Oct 22 – May 23 | 
 | GGG | Proffast | GGG | Proffast
TCCON |  |  | X | 
UoW EM27SUN/FO | X | X | X | X
NIWA EM27SUN | X | X | X | X
Uni Auckland EM27SUN | X |  | X | 
KIT EM27SUN |  |  | X | X
UoW IRcube | X | X | X | X

## Slide 18: Broken Hill Aircore campaign Sept/Oct 2022
2 x EM27/Sun with camtracker (COCCON, NIWA, Uni Auckland,)
1 x EM27 FTS with fibre-optic coupling to tracker (UoW)
1 x IRcube FTS with fibre-optic coupling to tracker (UoW)
7 days of measurements, including all aircore flight days

## Slide 19: Broken Hill aircore campaign Overview

## Slide 20: Broken Hill aircore campaign Raw retrievals, std AICF, ADCF (Pak et al 2022)

## Slide 21: AICF corrections relative to TCCON Wollongong
 | UOW-EM27FO | NIWA-EM27SUN | UoA- EM27SUN | IRcube
Xluft | 1.0019 | 0.9988 | 0.9982 | 1.0289
Xco2 | 1.0002 | 0.9980 | 0.9977 | 0.9969
Xch4 | 1.0009 | 0.9985 | 0.9980 | 1.0142

## Slide 22: Broken Hill aircore campaign AICF from Wollongong TCCON comparison

## Slide 23: GGG vs Proffast NIWA EM27SUN

## Slide 24: GGG vs Proffast NIWA EM27SUN

## Slide 25: Xair vs Xluft
Xair (Proffast)  1 / Xluft (GGG)
Xluft (GGG) = surface pressure – H2O column (in appropriate units)
O2 column/0.2095
Xair (PF) = O2 column/0.2095 + H2O column <= spectroscopic
surface pressure <= meteorological
Note: Frey et al (2019) COCCON paper (Proffit): Xair = Xluft

## Slide 26: GGG vs Proffast NIWA EM27SUN
Slope is equivalent to -10-15 sec time error

## Slide 27: Broken Hill aircore campaign NIWA EM27: TCCON AICF, -10 s time shift

## Slide 28: Intercomparison at Wollongong Oct 2022 – May 2023
4 EM27, 1 IRcube, TCCON
Benedikt Herkommer next talk: KIT travelling standard

## Slide 29: FO-coupled Ekotracker Pros and cons
Pros
  The Ekotracker is autonomous, robust, reliable and weatherproof
    Weather protection not required
  FTS is housed in controlled lab environment
    No big temperature changes
    Nicer for the FTS  - longer lifetime?
  Autonomous operation, no daily operator required
    Continuous measurements with direct sun
  Spectra and retrievals appear very similar to EM27SUN
    No spectrum artefacts
  Lower cost than Bruker camtracker + weather cover
Cons
  Glass/quartz lens/fibre cutoff => poorer results for CO and N2O
  Possible fibre optic artefacts, sensitivity to fibre alignment,
    Minor for EM27, worse for IRcube

## Slide 30: Next
Nail down entrance optics
Improve fibre alignment to centre on solar image on telescope
Extended range optics (CaF2 lenses, fibres?) to allow CO and N2O
Further characterise ADCF and AICF
  GGG and Proffast, ongoing
Trial at KIT beside standard COCCON FTS
Thank you

## Slide 31: Check AICF: EM27-FO – TCCON differences Xluft, XCO2, XCH4

## Slide 32: Effects of (mis)alignment of input optic ___EM27 ___IRcube
Rotate, move or swap fibre, limit intensity ~ noon each day

## Slide 33: EM27 ILS

## Slide 34: Wollongong side by side Oct 2022 – May 2023