## Slide 1: Log file for the retrieval of ILS at TAB
Working directories:
N2O: /data/Campaign/TAB/ilsFiles/lft145_n2o
HBr: /data/Campaign/TAB/ilsFiles/lft145_hbr
The working directories contain all the spectra and input files needed to retrieve the ILS with LINEFIT v145
Goal:
Participate in the N2O cell inter-comparison study of the NDACC-IRWG
To know the ILS of the updated instrument at Thule
Understand the important parameters in retrieval of the ILS and the information content

## Slide 2
N2O
Results presented here are from two N2O spectra taken on March 11 2016:
N2O#33_InSb_110316.0
N2O#33_InSb_110316.1
Two background spectra are used to create a normalized spectra to cancel out the N2O and CO2 from the lab:
N2O-NoCell_InSb_100316.0
N2O-NoCell_InSb_100316.1
The most important input parameters are (Following Frank’s input file):
3 micro windows: 2167.03 - 2185.2
2222.825 - 2223.019
2224.457 - 2224.715
Temperature: 291.9 K, column: 4.478e20, pressure: 0.8734 mbar
Maximal OPD: 180, Maximal inclination: 1.55e-3

## Slide 3
Example of the normalization using the three windows used in LINEFIT
N2O Spectrum: N2O#33_InSb_110316.1
Background: N2O-NoCell_InSb_100316.1
For the wide window we actually see a difference in the base line
If using the normalized spectra.
Note that the background spectrum was taken a day prior to the cell

## Slide 4
Results using the normalized spectra
Results using the non normalized spectra

## Slide 5
Results using the normalized spectra
Results using the non normalized spectra

## Slide 6
Results using the normalized spectra
Results using the non normalized spectra

## Slide 7
Remarks
There’s a significant improvement in the residual/RMS and in the modulation efficiency if the N2O spectra is normalized with the background. The improvement is significant for the long window where the baseline transmission is corrected.
The retrieved temperature is 296.2 K (initial Temp was 291.9 K).
A second run using the retrieved temperature does not change the retrieval parameters further.
The retrieval results using the other N2O spectrum are very similar as the ones presented here. However, the residuals are larger,
Recommendations: using a backgrounds spectra prior and after the N2O cell measurements.
Things to understand: Averaging kernels, how to improve the averaging kernels, test other windows?

## Slide 8
HBr
There are 5 spectra that can be used to retrieve the ILS:
HBr#60_InSb_040316.0,
HBr#60_InSb_040316.1  Better fits are obtained if used this spectra.
HBr#60_InSb_070316.0,
HBr#85_InSb_080316.0,
HBr#85_InSb_080316.1
Results presented here are from HBr#60_InSb_040316.1 taken on March 04 2016:
There are not background spectra for
The most important input parameters are 14 microwindows
Temperature: 293.3 (retrieved) K, column: 9.88e20 (fixed), pressure: 2.0 mbar
Maximal OPD: 257.0 (0.0035 cm-1), Maximal inclination: 1.37e-3 (APT: 1.15 mm / FOC: 418 mm)

## Slide 9
HBr

## Slide 10
Selected windows

## Slide 11
Remarks
The FWHM of the ILS obtained with HBr is lower (0.0028) than N2O cell (0.0038). The shape is similar
The modulation is closer to unity and the deviation shape is somehow similar as the N2O.
A background spectrum may reduce further the residuals.
Things to understand: Averaging kernels, how to improve the averaging kernels, test other windows?